|  | INTERNATIONAL TELECOMMUNICATION UNION |
| --- | --- |

CCITT	O.152

THE  INTERNATIONAL	(10/92)
TELEGRAPH  AND  TELEPHONE
CONSULTATIVE  COMMITTEE

SPECIFICATIONS  OF  MEASURING  EQUIPMENT

ERROR  PERFORMANCE  MEASURING
EQUIPMENT  FOR  BIT  RATES  OF
64  kbit/s  AND  N ´ 64  kbit/s

|  | Recommendation O.152 |
| --- | --- |

# FOREWORD

The CCITT (the International Telegraph and Telephone Consultative Committee) is a permanent organ of the International Telecommunication Union (ITU). CCITT is responsible for studying technical, operating and tariff questions and issuing Recommendations on them with a view to standardizing telecommunications on a worldwide basis.

The Plenary Assembly of CCITT which meets every four years, establishes the topics for study and approves Recommendations prepared by its Study Groups. The approval of Recommendations by the members of CCITT between Plenary Assemblies is covered by the procedure laid down in CCITT Resolution No. 2 (Melbourne, 1988).

Recommendation O.152 was revised by Study Group IV and was approved under the Resolution No. 2 procedure on the 5th of October 1992.

___________________

# CCITT  NOTE

In this Recommendation, the expression “Administration” is used for conciseness to indicate both a telecommunication administration and a recognized private operating agency.

ã  ITU  1993

All rights reserved. No part of this publication may be reproduced or utilized in any form or by any means, electronic or mechanical, including photocopying and microfilm, without permission in writing from the ITU.

Recommendation O.152

Recommendation O.152     (10/92)

ERROR  PERFORMANCE  MEASURING  EQUIPMENT
FOR  BIT  RATES  OF  64 kbit/s  AND  N  ´  64 kbit/sXE "ERROR  PERFORMANCE  MEASURING  EQUIPMENT
FOR  BIT  RATES  OF  64 kbit/s  AND  N  ´  64 kbit/s"§

(Published 1984; revised 1988, 1992)

## Abstract

Defines the requirements for an equipment to measure bit-error performance at 64 kbit/s and N  ´ 64 kbit/s.

## Keywords

–	digital error detector;

–	digital pattern generator;

–	error performance measurement;

–	 measurement;

– 	tester.

PREAMBLE

The requirements for the characteristics of a bit-error performance measuring equipment which are described below must be adhered to in order to ensure compatibility between equipment produced by different manufacturers.

# 1	General

The equipment is designed to measure the bit-error performance of digital paths (operating at 64 kbit/s and N  ´ 64 kbit/s) by the direct comparison of a pseudo-random test pattern with locally generated test pattern identical to the transmitted test pattern.

# 2	Test patterns

The following patterns are recommended (see Recommendation O.150 [8], for further details).

## 2.1	Pseudo-random pattern of 211 – 1 (2047 bit) pattern length

This pattern is primarily intended for error and jitter measurements on circuits operating at bit rates of 64 kbit/s and N  ´ 64 kbit/s.

The pattern may be generated in an eleven-stage shift register whose 9th and 11th stage outputs are added in a modulo-two addition stage, and the result is fed back to the input of the first stage.

Number of shift register stages		11

Length of pseudo-random sequence	211 – 1 = 2047 bits

Longest sequence of ZEROs		10 (non-inverted signal)

Note 1 – In the case of international testing where the measurement includes systems based on 1544 kbit/s, it is necessary to modify the test sequence in such a way that more than seven consecutive ZEROs are avoided. This is achieved by forcing the output bit to ONE whenever the previous 7 bits of the sequence are all ZEROs.

Note 2 – It is recommended to use the test pattern of 2047 bit length also at other bit rates in the range of 48 kbit/s to 168 kbit/s.

Note 3 – When performing measurements at a bit rate of N  ´ 64 kbit/s, consecutive 8-bit blocks of the test pattern shall be transmitted in consecutive time slots. In this case, the beginning of the test pattern need not to be related to the beginning of a frame.

Note 4 – Whether N can be any number between 1 and 31 requires further study.

## 2.2	Fixed patterns (optional)

Fixed patterns of all ONEs (...1111...) and alternating ONEs and ZEROs (...1010...) may be provided.

## 2.3	Loss of sequence synchronizationXE "Loss of sequence synchronization"§

Sequence synchronization shall be considered to be lost and resynchronization shall be started if:

a)	the bit-error ratio is  ³  0.2 during an integration interval of 1 second; or

b)	it can be unambiguously identified that the test pattern and the reference pattern are out of phase.

Note – One method to recognize the out-of-phase condition is the evaluation of the error pattern resulting from the bit-by-bit comparison. If the error pattern has the same structure as the pseudo-random test pattern, then the out-of-phase condition can be recognized.

This sub-section requires further study.

# 3	Bit rate

Bit rate in accordance with § 1 of Recommendation G.703 [1] and Recommendation V.36 [2] of 64 kbit/s:

a)	bit rate tolerance (Recommendation G.703 [1]):  ± 100 × 10–6,

b)	bit rate tolerance (Recommendation V.36 [2]), optional:  ± 50 × 10–6.

# 4	Interfaces

The interface characteristics (impedances, levels, codes, etc.) should be in accordance with Recommendations G.703 [1], I.430 [7] (optional) and V.11 [3] (optional).

In addition to providing for terminated measurements, the measuring set shall also be capable of monitoring at protected test points on digital equipment. Therefore, a high impedance and/or additional gain must be provided to compensate for the loss at monitoring points already provided on some equipments.

## 4.1	Interfaces corresponding to Recommendation G.703 [1]

Three interfaces shall be provided:

a)	a codirectional interface in accordance with § 1.2.1 of Recommendation G.703 [1];

b)	a centralized clock interface in accordance with § 1.2.2 of Recommendation G.703 [1];

c)	a contradirectional interface in accordance with § 1.2.3 of Recommendation G.703 [1].

## 4.2	Method of clock synchronizationXE "Method of clock synchronization"§

The following modes of synchronization shall be selectable:

a)	Lock the digital generator clock rate to that at the input of the receive side of the measuring set (for the codirectional interface).

b)	Allow the generator clock to free run within the overall allowed frequency tolerances.

c)	Lock the digital generator clock rate to an external clock signal. (Configuration of input for external clock in accordance with Recommendation G.703 [1].)

## 4.3	Interface corresponding to Recommendation I.430 [7]

For further study. This study should include means for obtaining access to the individual 64 kbit/s channels at the S and T interface points.

## 4.4	Interface corresponding to Recommendation V.11 [3]

As an option, an interface in accordance with Recommendation V.11 [3] shall be provided.

# 5	Bit-error ratio measuring rangeXE "Bit-error ratio measuring range"§

The receiving equipment of the set should be capable of measuring bit-error ratios in the range 10–2 to 10–7. The measurement time should be sufficiently long to achieve accurate measurements. In addition, it should be possible to measure bit-error ratios smaller than 10–7. This can be achieved by providing the capability to count cumulative errors.

# 6	Block-error ratio measurementsXE "Block-error ratio measurements"§

Optionally, the instrument should be capable of performing block-error measurements in addition to the bit-error measurements. If provided, it should be possible to measure block-error ratios in the range 100 to 10–5 when using the pseudo-random test pattern with a block length of 2047 bits.

# 7	Mode of operation

The mode of operation should be such that the signal to be tested is first converted into a unipolar (binary) signal in the error measuring instrument and subsequently the bit comparison is made also with a reference signal in binary form.

# 8	Error evaluationXE "Error evaluation"§

## 8.1	Measurement of errored time intervalsXE "Measurement of errored time intervals"§

The instrument shall be capable of detecting errored seconds and other errored or error-free time intervals as defined by Recommendation G.821 [4]. The number of errored or error-free time intervals in a selectable observation period from 1 minute to 24 hours, or continuous, shall be counted and displayed.

For this measurement, the error detection circuits of the instrument shall be controlled by an internal timer which sets intervals of equal length and which operates independently of the occurrence of errors.

## 8.2	Measurement of short-term mean error ratioXE "Measurement of short-term mean error ratio"§

8.2.1	It shall be possible to record the time intervals as defined in Recommendation G.821 [4], during which the bit-error ratio is less than 1 × 10–6.

8.2.2	It shall be possible to record the one-second intervals during which the bit-error ratio is less than 1 × 10–3.

# 9	Recording of measurement resultsXE "Recording of measurement results"§

As an option an interface shall be provided which allows connecting external equipment for further processing the measuring results.

The interface shall comply with Recommendation V.24 [5] or the interface bus according to IEC Publication 625 [6].

# 10	Operating environment

The electrical performance requirements shall be met when operating at the climatic conditions as specified in § 2.1 of Recommendation O.3 [9].

References

[1]	CCITT Recommendation G.703 – Physical/electrical characteristics of hierarchical digital interfaces.

[2]	CCITT Recommendation V.36 – Modems for synchronous data transmission using 60-108 kHz group band circuits.

[3]	CCITT Recommendation V.11 – Electrical characteristics for balanced double-current interchange circuits for general use with integrated circuit equipment in the field of data communications.

[4]	CCITT Recommendation G.821 – Error performance on an international digital connection forming part of an integrated services digital network.

[5]	CCITT Recommendation V.24 – List of definitions for interchange circuits between data terminal equipment and data circuit-terminating equipment.

[6]	IEC Publication 625 – An Interface system for programmable measuring instruments (byte serial, bit parallel).

[7]	CCITT Recommendation I.430 – Basic user-network interface-Layer/Specification.

[8]	CCITT Recommendation O.150 – Digital test patterns for performance measurements on digital transmission equipment.

[9]	CCITT Recommendation O.3 – Climatic conditions and relevant tests for measuring equipment.