
	[image: image1.wmf]
	INTERNATIONAL  TELECOMMUNICATION  UNION

	
	

	ITU-T
	O.172

	TELECOMMUNICATION
STANDARDIZATION  SECTOR
OF  ITU
	(03/2001)   

	
	SERIES O: SPECIFICATIONS OF MEASURING EQUIPMENT

Equipment for the measurement of digital and analogue/digital parameters



	
	Jitter and wander measuring equipment for digital systems which are based on the synchronous digital hierarchy (SDH)

	
	ITU‑T  Recommendation  O.172

(Formerly  CCITT  Recommendation)


ITU-T O-SERIES  RECOMMENDATIONS

SPECIFICATIONS OF MEASURING EQUIPMENT

	
	

	General
	O.1–O.9

	Maintenance access
	O.10–O.19

	Automatic and semi-automatic measuring systems
	O.20–O.39

	Equipment for the measurement of analogue parameters
	O.40–O.129

	Equipment for the measurement of digital and analogue/digital parameters
	O.130–O.199

	
	


For further details, please refer to the list of ITU-T Recommendations.

	ITU-T Recommendation O.172

Jitter and wander measuring equipment for digital systems
which are based on the synchronous digital hierarchy (SDH)



	Summary

This Recommendation specifies instrumentation that is used to generate and measure jitter and wander in digital systems based on the SDH. Measurement requirements for both SDH line interfaces and SDH tributary interfaces operating at PDH bit rates are addressed in this Recommendation.

The requirements for the characteristics of the jitter and wander measuring equipment that are specified in this Recommendation must be adhered to in order to ensure consistency of results between equipment produced by different manufacturers.


	Source

ITU‑T Recommendation O.172 was revised by ITU‑T Study Group 4 (2001‑2004) and approved under the WTSA Resolution 1 procedure on 15 March 2001.


	Keywords

Input jitter tolerance, Input wander tolerance, Jitter generation, Jitter measurement, Jitter transfer function, Output jitter, Output wander, Phase transients, Pointer jitter, Pointer sequence generation, Wander generation, Wander measurement, Wander noise transfer.


FOREWORD

The International Telecommunication Union (ITU) is the United Nations specialized agency in the field of tele​com​mu​ni​ca​tions. The ITU Telecommunication Standardization Sector (ITU-T) is a permanent organ of ITU. ITU-T is responsible for studying technical, operating and tariff questions and issuing Recommendations on them with a view to standardizing telecommunications on a worldwide basis.

The World Telecommunication Standardization Assembly (WTSA), which meets every four years, establishes the topics for study by the ITU‑T study groups which, in turn, produce Recommendations on these topics.

The approval of ITU-T Recommendations is covered by the procedure laid down in WTSA Resolution 1.

In some areas of information technology which fall within ITU-T's purview, the necessary standards are prepared on a collaborative basis with ISO and IEC.

NOTE

In this Recommendation, the expression "Administration" is used for conciseness to indicate both a telecommunication administration and a recognized operating agency.

INTELLECTUAL PROPERTY RIGHTS
ITU draws attention to the possibility that the practice or implementation of this Recommendation may involve the use of a claimed Intellectual Property Right. ITU takes no position concerning the evidence, validity or applicability of claimed Intellectual Property Rights, whether asserted by ITU members or others outside of the Recommendation development process.

As of the date of approval of this Recommendation, ITU had not received notice of intellectual property, protected by patents, which may be required to implement this Recommendation. However, implementors are cautioned that this may not represent the latest information and are therefore strongly urged to consult the TSB patent database.

(  ITU  2001

All rights reserved. No part of this publication may be reproduced or utilized in any form or by any means, electronic or mechanical, including photocopying and microfilm, without permission in writing from ITU.

CONTENTS


Page

1
Scope

1
2
References

1
2.1
Normative references

1
2.2
Informative references

2
3
Definitions

2
4
Abbreviations

2
5
Conventions

3
6
Functional block diagram

3
7
Interfaces

4
7.1
Optical interfaces

4
7.2
Electrical interfaces

5
7.3
External reference clock input

5
7.4
Input interface sensitivity

5
8
Jitter/wander generation function

5
8.1
Modulation source

6
8.2
Clock generator

6
8.2.1
Accuracy of the clock generator

6
8.3
Digital test pattern generator

6
8.3.1
Digital test patterns

6
8.3.2
Digital test patterns for SDH tributary signals

6
8.4
Pointer sequence generator

7
8.4.1
Pointer test sequence generation capability

7
8.5
Minimum jitter/wander generation capability

8
8.6
Generation accuracy

10
8.6.1
Phase amplitude error

10
8.6.2
Phase slope error

10
8.6.3
Intrinsic jitter/wander of generation function

11
9
Jitter measurement function

12
9.1
Reference timing signal

12
9.2
Measurement capabilities

12
9.2.1
Measurement range

12
9.2.2
Selectable threshold

13
9.2.3
Measurement of RMS jitter

13
9.2.4
Input phase tolerance for SDH tributary signals

13

Page

9.3
Measurement bandwidths

14
9.3.1
Frequency response of jitter measurement function for SDH line signals

15
9.3.2
Frequency response of jitter measurement function for SDH tributary signals

15
9.4
Measurement accuracy

16
9.4.1
Measurement result accuracy

16
9.4.2
Fixed error of SDH line jitter measurements

16
9.4.3
Fixed error of SDH tributary jitter measurements

17
9.4.4
Variable error of SDH tributary jitter measurements

17
9.4.5
Digital test signal-dependent error

18
9.5
Additional facilities

19
9.5.1
Analogue output

19
10
Wander measurement function

19
10.1
Reference timing signal

19
10.2
Measurement of TIE (Time Interval Error)

19
10.2.1
Sampling interval

19
10.2.2
Measurement bandwidth

19
10.2.3
Measurement range

19
10.2.4
Measurement result accuracy

20
10.3
Measurement of transient TIE (Time Interval Error)

20
10.3.1
Sampling Interval

20
10.3.2
Measurement Bandwidth

20
10.3.3
Measurement Range

20
10.3.4
Measurement Result Accuracy

21
10.4
Measurement of MTIE (Maximum Time Interval Error)

21
10.4.1
Measurement and observation interval ranges

21
10.4.2
Calculation algorithm accuracy

21
10.4.3
Measurement result accuracy

22
10.5
Measurement of TDEV (Time Deviation)

22
10.5.1
Measurement and observation interval ranges

22
10.5.2
Calculation algorithm accuracy

22
10.5.3
Measurement result accuracy

23
10.6
Measurement of Frequency Offset

23
10.6.1
Measurement range

24
10.6.2
Calculation algorithm accuracy

24
10.6.3
Measurement result accuracy

24
10.7
Measurement of Frequency Drift Rate

25
10.7.1
Measurement range

25
10.7.2
Calculation algorithm accuracy

25

Page

10.7.3
Measurement result accuracy

26
11
TDEV wander noise generation function

26
12
Operating environment

27
Annex A ( Structured test signals for the measurement of jitter

27
A.1
Introduction

27
A.1.1
Payload test conditions

27
A.1.2
SDH overhead byte conditions

27
A.2
Test signal structure for STM‑N signals

27
A.2.1
STM‑1 signal

27
A.2.2
STM‑N signal (N ( 4)

28
A.2.3
STM‑0 signal

28
A.3
Test signal structure for concatenated STM‑N signals

28
A.3.1
STM‑N signal (N ( 4)

28
Annex B ( Definition of band-limited peak-to-peak phase slope error

29
Annex C ( Specification of distribution for TDEV wander noise generation

30
Appendix I ( Guidelines concerning the measurement of jitter in SDH systems

30
Appendix II ( Guidelines concerning the measurement of wander in SDH systems

31
II.1
Wander measurements

31
II.1.1
General considerations on wander measurement configurations

31
II.1.2
Synchronized wander measurements

31
II.1.3
Non-synchronized wander measurements

32
II.2
Clock stability measurements

32
Appendix III ( Guidelines concerning the generation of pointer test sequences

33
Appendix IV ( Total jitter measurement function response

34
IV.1
Introduction

34
IV.2
Measurement filter parameters

34
IV.3
Mask limits for high-pass measurement filter response

34
IV.3.1
SDH tributary jitter measurement high-pass filter

35
Appendix V ( Verification of MTIE and TDEV calculation algorithms

36
V.1
TIE noise source functional description

36
V.2
First example of TIE noise generator

36
V.3
Second example of TIE noise generator

36
Introduction

The timing and synchronization performance of SDH networks and SDH network equipment elements is specified in a number of ITU‑T Recommendations, using jitter and wander parameters. This Recommendation specifies the various characteristics of jitter/wander measuring equipment, which are needed in order to support the requirements of these ITU‑T Recommendations and to perform other test and measuring tasks.

This Recommendation has been developed to ensure maximum compatibility with the relevant SDH network and equipment measurement requirements, whilst maintaining backwards compatibility with the associated PDH test equipment requirements of ITU‑T O.171 [17] where possible.

While functional and characteristic requirements are given for the measuring equipment, the realization of the equipment configuration is not covered and should be given careful consideration by the designer and user. In particular, it is not required that all features described in this Recommendation shall be provided in one piece of equipment. Users may select those functions which correspond best to their applications.

ITU-T Recommendation O.172

Jitter and wander measuring equipment for digital systems 
which are based on the synchronous digital hierarchy (SDH)

1
Scope

This Recommendation specifies test instrumentation that is used to generate and measure timing jitter and synchronization wander in digital systems based on the Synchronous Digital Hierarchy (SDH).

This Recommendation also specifies requirements for the measurement of SDH tributaries operating at PDH bit rates.

The test instrumentation consists principally of a jitter/wander measurement function and a jitter/wander generation function. Measurements can be performed at the physical layer of SDH systems. A bit error rate test set may also be required for certain types of measurements; this may be part of the same instrumentation or it may be physically separate.

Test instrumentation for the generation and measurement of jitter and wander in digital systems based on the Plesiochronous Digital Hierarchy (PDH) is specified in ITU‑T O.171 [17].

It is recommended that ITU‑T G.783 [5], ITU‑T G.812 [8], ITU‑T G.813 [9], ITU‑T G.825 [12] and ITU‑T G.958 [14] be read in conjunction with this Recommendation.

2
References

The following ITU-T Recommendations and other references contain provisions which, through reference in this text, constitute provisions of this Recommendation. At the time of publication, the editions indicated were valid. All Recommendations and other references are subject to revision; all users of this Recommendation are therefore encouraged to investigate the possibility of applying the most recent edition of the Recommendations and other references listed below. A list of the currently valid ITU-T Recommendations is regularly published.

2.1
Normative references

[1]
ITU‑T G.691 (2000), Optical interfaces for single-channel STM-64, STM-256 and other SDH systems with optical amplifiers.

[2]
ITU-T G.703 (1998), Physical/electrical characteristics of hierarchical digital interfaces.
[3]
ITU‑T G.707/Y.1322 (2000), Network node interface for the synchronous digital hierarchy (SDH).
[4]
ITU‑T G.772 (1993), Protected monitoring points provided on digital transmission systems.
[5]
ITU‑T G.783 (2000), Characteristics of synchronous digital hierarchy (SDH) equipment functional blocks.
[6]
ITU‑T G.810 (1996), Definitions and terminology for synchronization networks.
[7]
ITU‑T G.811 (1997), Timing characteristics of primary reference clocks.
[8]
ITU‑T G.812 (1998), Timing requirements of slave clocks suitable for use as node clocks in synchronization networks.
[9]
ITU‑T G.813 (1996), Timing characteristics of SDH equipment slave clocks (SEC).
[10]
ITU‑T G.823 (2000), The control of jitter and wander within digital networks which are based on the 2048 kbit/s hierarchy.
[11]
ITU‑T G.824 (2000), The control of jitter and wander within digital networks which are based on the 1544 kbit/s hierarchy.
[12]
ITU‑T G.825 (2000), The control of jitter and wander within digital networks which are based on the synchronous digital hierarchy (SDH).
[13]
ITU‑T G.957 (1999), Optical interfaces for equipments and systems relating to the synchronous digital hierarchy.
[14]
ITU‑T G.958 (1994), Digital line systems based on the synchronous digital hierarchy for use on optical fibre cables.
[15]
ITU-T O.3 (1992), Climatic conditions and relevant tests for measuring equipment.
[16]
ITU‑T O.150 (1996), General requirements for instrumentation for performance measurements on digital transmission equipment.
[17]
ITU‑T O.171 (1997), Timing jitter and wander measuring equipment for digital systems which are based on the plesiochronous digital hierarchy (PDH).
[18]
ITU‑T O.181 (1996), Equipment to assess error performance on STM‑N interfaces.

[19]
ITU‑R F.750-3 (1997), Architectures and functional aspects of radio-relay systems for SDH-based networks.
2.2
Informative references

[20]
ANSI T1.105.03, 1994, Telecommunications – Synchronous Optical Network (SONET) – Jitter at Network Interfaces.
[21]
ANSI T1.105.06, 1996, Telecommunications – Synchronous Optical Network (SONET) – Physical Layer Specification.
3
Definitions

This Recommendation defines the following terms (refer to ITU-T G.810 [6]):

3.1
(timing) jitter: The short-term variations of the significant instances of a digital signal from their ideal positions in time (where "short-term" implies that these variations are of frequency greater than or equal to 10 Hz).

3.2
wander: The long-term variations of the significant instances of a digital signal from their ideal position in time (where "long-term" implies that these variations are of frequency less than 10 Hz).

3.3
time interval error (function): The difference between the measure of a time interval as provided by a clock and the measure of that same time interval as provided by a reference clock.

It may be useful to note that ITU-T G.810 [6] provides additional definitions and abbreviations used in timing and synchronization Recommendations. It also provides background information on the need to limit phase variation and the impairments on digital systems.

4
Abbreviations

This Recommendation uses the following abbreviations:

AU‑n
Administrative Unit, level n

CMI
Coded Mark Inversion

MTIE
Maximum Time Interval Error

NRZ
Non Return to Zero

PDH
Plesiochronous Digital Hierarchy

PJE
Pointer Justification Event

PLL
Phase-Locked Loop

ppm
parts per million

PRBS
Pseudo Random Binary Sequence

RMS
Root-Mean-Square

SDH
Synchronous Digital Hierarchy

STM‑N
Synchronous Transport Module, level N

TDEV
Time Deviation

TIE
Time Interval Error

TSS
Test Signal Structure

TU‑m
Tributary Unit, level m

UI
Unit Interval

UIpp
Unit Interval, peak-to-peak

5
Conventions

For the purposes of this Recommendation, the following conventions are adopted:

a)
Particular interface signals used are denoted either by their standardized signal formats, e.g. STM‑1 or by their bit rate, e.g. 139 264 kbit/s. The default physical format of SDH interfaces is considered to be optical and the default physical format of PDH interfaces is considered to be electrical.

b)
Where the electrical form of an SDH interface is specifically used in this Recommendation, the interface is denoted by "e", e.g. STM‑1e (refer also to clause 7).

c)
Particular interface signals used may be categorized either as SDH line interfaces, or as SDH tributary interfaces. In this Recommendation, SDH line interfaces refer to those which support STM‑N signals, whilst SDH tributary interfaces refer to those which support signals operating at PDH bit rates.

6
Functional block diagram

Figure 1 shows the block diagram of the instrumentation in general form, identifying the main functions that are addressed in this Recommendation. The figure does not describe a specific implementation.


[image: image2.wmf]T0409200-98

External

reference

clock input

Modulation

source

Clock

generator

Pointer

sequence

generator*

Digital test

pattern

generator

Interfaces

Jitter/wander

generation

function

Test

signal

output

Reference

timing

signal

Phase

detector

Measurement

filter

Jitter

measurement 

function

Jitter

result

* SDH only

Interfaces

Test

signal

input

Measurement

filter

Phase

detector

MTIE

calculation

TIE

result

MTIE

result

External

reference

clock input

Reference

timing

signal

TDEV

calculation

Wander

measurement

function

TDEV

result


Figure 1/O.172 – Functional block diagram for jitter and wander test set

7
Interfaces

7.1
Optical interfaces

The instrumentation shall be capable of operating at one or more of the following bit rates and corresponding optical interface characteristics as defined in the appropriate clauses of ITU‑T G.957 [13] and ITU‑T G.691 [1] and also ANSI T1.105.06 [21] (for 51 840 kbit/s interfaces).

–
STM-0
51 840 kbit/s

–
STM-1
155 520 kbit/s

–
STM-4
622 080 kbit/s

–
STM-16
2 488 320 kbit/s

–
STM-64
9 953 280 kbit/s

–
STM-256
39 813 120 kbit/s

7.2
Electrical interfaces

The instrumentation shall be capable of operating at one or more of the following bit rates and corresponding electrical interface characteristics as defined in the appropriate clauses of ITU‑T G.703 [2] and also ITU‑R F.750-3 [19] (for 51 840 kbit/s electrical interfaces). However, for all bit rates the signal applied to the input of the jitter/wander measuring circuit shall be a nominal rectangular pulse. Other signal shapes may produce inter-symbol interference – which cannot be corrected by simple line equalization – thus affecting measurement accuracy.

–
1544 kbit/s

–
2048 kbit/s

–
6312 kbit/s

–
34 368 kbit/s

–
44 736 kbit/s

–
51 840 kbit/s,
STM‑0e

–
139 264 kbit/s

–
155 520 kbit/s,
STM‑1e

The jitter/wander measurement function input and jitter/wander generation function output ports shall have a return loss as specified in the appropriate clauses of ITU‑T G.703 [2].

As an option, the jitter/wander measurement function shall be capable of measuring jitter/wander at a clock output port when such an access is provided on digital equipment.

7.3
External reference clock input

The measuring equipment shall accept data signals at bit rates of 1544 kbit/s or 2048 kbit/s as a reference. If 2048 kbit/s can be accepted, the equipment shall also accept a clock signal at 2048 kHz as a reference. The characteristics of the clock signals shall be in accordance with ITU‑T G.703 [2].

7.4
Input interface sensitivity

The jitter/wander measurement function is required to operate satisfactorily under the following input conditions:

a)
the specification for equipment optical interfaces defined in ITU‑T G.957 [13] and G.691 [1];

b)
the specification for equipment electrical interfaces defined in ITU‑T G.703 [2];

c)
protected monitoring points as defined in ITU‑T G.772 [4].

8
Jitter/wander generation function

Tests of digital equipment may be made with either a jittered, wandered or a non-jittered/wandered digital signal. This will require the digital test pattern generator, clock generator and modulation source shown in Figure 1.

8.1
Modulation source

The modulation source, required to perform tests in conformance with relevant Recommendations, may be provided within the clock generator and/or digital test pattern generator or it may be provided separately. In this Recommendation, the modulation source is defined to be sinusoidal. However, other stimuli may be required for certain tests.

8.2
Clock generator

It shall be possible to phase-modulate the clock generator from the modulation source and to indicate the peak-to-peak phase deviation of the modulated signal.

The generated peak-to-peak jitter/wander and the modulating frequencies shall meet the minimum requirements of Table 3 and Figure 3.

If the output interfaces for the modulated clock signal and/or the external timing reference signal are provided, the minimum amplitude shall be 1 Volt peak-to-peak into 75 ( or 0.25 Volt peak-to-peak into 50 (.

8.2.1
Accuracy of the clock generator

The frequency deviation of the internal clock signal from its nominal value shall be less than:


(4.6 ppm

As an option, the clock generator may provide adjustable frequency offset of sufficient magnitude to facilitate testing across the clock tolerance range of the equipment-under-test, e.g. (10 ppm to (100 ppm, as defined for the various bit rates in ITU‑T G.703 [2] and ITU‑T G.813 [9].

It shall be possible to phase-lock the generation function to an external reference clock source of arbitrary accuracy; refer also to 7.3.

8.3
Digital test pattern generator

The jitter/wander measurement function will normally be used in conjunction with any suitable digital test pattern generator providing the following facilities.

8.3.1
Digital test patterns

The digital test pattern generator shall be capable of providing one or more of the following signals, for use at STM‑N bit rates:

a)
framed SDH signals in accordance with ITU‑T G.707 [3];

b)
structured test signals in accordance with ITU‑T O.181 [18], dependant on the type of network element to be tested;

c)
structured test signals defined in Annex A.

8.3.2
Digital test patterns for SDH tributary signals

The test sequence generator shall be capable of providing the following signals:

For use at bit rates of 2048 kbit/s, 6312 kbit/s, and 44 736 kbit/s, a pseudo random test sequence of 215 – 1 length corresponding to 5.3/O.150 [16].

For use at bit rates of 1544 kbit/s, 6312 kbit/s, and 44 736 kbit/s, a pseudo random test sequence of 220 – 1 bit length corresponding to 5.5/O.150 [16].

For use at bit rates of 34 368 kbit/s and 139 264 kbit/s, a pseudo random test sequence of 223 – 1 bit length corresponding to 5.6/O.150 [16].

For use at all bit rates, a 10001000 repetitive test sequence.

As an option and for use at all bit rates:

a)
two freely programmable 8-bit test sequences capable of being alternated at a low rate (e.g. from 10 Hz to 100 Hz);

b)
a freely programmable 16-bit test sequence.

8.4
Pointer sequence generator

In order to test the effect of pointer justification events (PJEs) on SDH desynchronizer equipment, the digital test pattern generator shall be capable of generating complete test sequences of pointer justifications in accordance with the appropriate clauses of ITU‑T G.783 [5].

As an option, the equipment may provide additional PJE control functions to facilitate further pointer jitter testing.

Appendix III provides further information regarding the test set configuration and capability for testing using pointer sequences.

8.4.1
Pointer test sequence generation capability

The equipment shall provide a minimum set of pointer test sequences applicable at particular SDH tributary bit rates, denoted by "X" in Table 1, in accordance with clause 10/G.783 [5].

For particular SDH tributary bit rates, the pointer adjustments shall be applied either to the AU‑n or the TU‑m pointers, as shown in Table 1. The direction or polarity of the pointer test sequence shall be selectable between incrementing or decrementing pointer values.

The test procedure specified in clause 15/G.783 [5] shall be followed.

Table 1/O.172 – G.783 pointer test sequence description

	G.783 pointer test sequence
	SDH tributary bit rate (kbit/s) and SDH container

	
	1544
	2048
	6312
	34 368
	44 736
	139 264

	ID
	Description
	TU‑11
	TU‑12
	TU‑2
	TU‑3
	AU‑3
	AU‑4

	a
	Single Alternating
	
	X
	(Note)
	X
	
	X

	b
	Regular ( Double
	
	X
	
	X
	
	X

	c
	Regular ( Missing
	
	X
	
	X
	
	X

	d
	Double Alternating
	
	
	
	X
	
	X

	e
	Single
	X
	
	
	
	X
	

	f
	Burst
	
	
	
	
	X
	

	g1
	Periodic 87-3
	
	
	
	
	X
	X

	g2
	Periodic 87-3 with Add
	
	
	
	
	X
	X

	g3
	Periodic 87-3 with Cancel
	
	
	
	
	X
	X

	h1
	Periodic
	X
	
	
	
	X
	

	h2
	Periodic with Add
	X
	
	
	
	X
	

	h3
	Periodic with Cancel
	X
	
	
	
	X
	

	i
	Phase Transient
	
	
	
	
	X
	

	j1
	Periodic 26-1
	X
	
	
	
	
	

	j2
	Periodic 26-1 with Add
	X
	
	
	
	
	

	j3
	Periodic 26-1 with Cancel
	X
	
	
	
	
	

	NOTE – Pointer test sequences for 6312 kbit/s, TU-2 are for further study.


The time intervals between PJEs within a test sequence depend on the particular sequence and the bit rate of the SDH tributary under test. ITU‑T G.783 [5] shall be consulted for precise detail of the pointer sequences applicable to a particular tributary bit rate. Table 2 specifies the time intervals that shall be provided between PJEs for particular SDH tributary bit rates, and which are illustrated in a generic fashion in Figure 2. Time intervals T1 and T2 may be adjustable to values greater than the minimum shown in Table 2.

Table 2/O.172 – Time intervals between G.783 pointer justification events

	Time Interval
	SDH tributary bit rate (kbit/s)

	
	1544
	2048
	6312
	34 368
	44 736
	139 264

	T1 (minimum)
	30 s
	10 s
	(Note)
	10 s
	30 s
	10 s

	T2 (minimum)
	1 s
	750 ms
	(Note)
	(Note)
	34 ms
	(Note)

	T3
	2 ms
	2 ms
	(Note)
	(Note)
	0.5 ms
	(Note)

	NOTE – Value is for further study in ITU-T G.783 [5].



[image: image3.wmf]T0409210-98

T3

T2

T1


Figure 2/O.172 – Generic G.783 pointer sequence, illustrating the time intervals of Table 1

8.5
Minimum jitter/wander generation capability

The jitter/wander amplitude/frequency characteristic of the generation function shall meet the minimum requirements of Figure 3 and Table 3 for SDH line signals or Table 3a for SDH tributary signals.

	Table 3/O.172 – Minimum amplitude of adjustable generated jitter/wander amplitude
versus jitter/wander frequency for SDH line signals

	Signal
	Minimum peak-to-peak 
jitter/wander amplitude (UIpp)
	Jitter/wander frequency breakpoints (Hz)

	
	A0
	A1
	A2
	A3
	A4
	f0
	f12
	f11
	f10
	f9
	f8
	f1
	f2
	f3
	f4

	STM‑0e,
STM‑0
	*
	*
	20
	2
	0.2
	*
	*
	*
	*
	10
	30
	300
	2 k
	20 k
	400 k

	STM‑1e,
STM-1
	3 600
	400
	

50
	2
	0.2
	12 (
	178 (
	1.6 m
	15.6 m
	125 m
	19.3
	500
	6.5 k
	65 k
	1.3 M

	STM‑4
	14 400
	1 600
	

200
	2
	0.2
	12 (
	178 (
	1.6 m
	15.6 m
	125 m
	9.65
	1 k
	25 k
	250 k
	5 M

	STM‑16
	57 600
	6 400
	

800
	2
	0.2
	12 (
	178 (
	1.6 m
	15.6 m
	125 m
	12.1
	5 k
	100 k
	1 M
	20 M

	STM‑64
	230 400
	25 600
	

3 200
	2
	0.2
	12 (
	178 (
	1.6 m
	15.6 m
	125 m
	12.1
	20 k
	400 k
	4 M
	80 M

	STM-256
	FFS
	FFS
	FFS
	FFS
	FFS
	FFS
	FFS
	FFS
	FFS
	FFS
	FFS
	FFS
	FFS
	FFS
	FFS

	NOTE 1 – Values denoted by "*" are not defined.

NOTE 2 – Values are based on the requirements of Table 2/G.825 [12] and Figure 2/G.825 [12].

NOTE 3 – Values for STM‑0 are based on the requirements of ANSI T1.105.03 [20].

NOTE 4 – FFS denotes that the value is for further study.


Table 3a/O.172 – Minimum amplitude of adjustable generated jitter/wander 
amplitude versus jitter/wander frequency for SDH tributaries

	Signal
(kbit/s)
	Minimum peak-to-peak 
jitter/wander amplitude (UIpp)
	Jitter/wander frequency breakpoints (Hz)

	
	A0
	A1
	A2
	A3
	A4
	f0
	f12
	f11
	f10
	f9
	f8
	f1
	f2
	f3
	f4

	1544
	40
	*
	
20
	10
	0.5
	12 µ
	
50 m
	*
	*
	
100 m
	5
	
10
	
400
	
8 k
	
40 k

	2048
	50
	*
	
30
	10
	0.5
	12 µ
	
12 m
	*
	*
	
20 m
	3.3
	
10
	
900
	
18 k
	
100 k

	6312
	150
	*
	
50
	10
	0.5
	12 µ
	
3 m
	*
	*
	
10 m
	0.4
	
2
	
1600
	
32 k
	
60 k

	34 368
	200
	*
	
50
	10
	0.5
	10 m
	
50 m
	*
	*
	
200 m
	20
	
100
	
1000
	
20 k
	
800 k

	44 736
	
1000
	*
	
100
	10
	0.5
	12 µ
	
11 m
	*
	*
	
110 m
	2.19
	
21.9
	5000
	
100 k
	
400 k

	139 264
	800
	*
	
200
	10
	0.5
	10 m
	
33 m
	*
	*
	
130 m
	5
	
100
	
500
	
10 k
	3500 k

	NOTE 1 – Values denoted by "*" are not defined.

NOTE 2 – These requirements are based on consideration of ITU‑T G.823 [10] and ITU‑T G.824 [11].



[image: image4.wmf]T0409220-98

A

0

A

1

A

2

A

3

A

4

f

0

f

12

f

11

f

10

f

9

f

8

f

1

f

2

f

3

f

4

Peak-to-peak

jitter/wander amplitude

(log scale)

Frequency (log scale)


Figure 3/O.172 – Generated jitter/wander amplitude
versus jitter/wander frequency

8.6
Generation accuracy

The test signal source shall be compatible with the jitter/wander measurement function in such a way that the overall measuring accuracy is not substantially deteriorated. The generation accuracy may be increased by measuring the jitter/wander applied to the unit under test using a corresponding jitter/wander measuring device.

The generating accuracy of the jitter/wander generation function is dependent upon several factors such as fixed intrinsic error, setting resolution, distortion and frequency response error. In addition, there is an error that is a function of the actual setting.

8.6.1
Phase amplitude error

The amplitude error of sinusoidal jitter/wander generation shall be less than:


(Q% of setting (0.02 UIpp

where Q is a variable error specified in Table 4 for SDH line signals and in Table 4a for SDH tributary signals. The frequencies f0, f1, f4 and f9 used in Tables 4 and 4a are defined in Tables 3 and 3a.

NOTE – This specification excludes any wideband intrinsic jitter/wander components.

8.6.2
Phase slope error

The band-limited peak-to-peak phase slope error in UI/s shall be less than:




[image: image5.wmf]2

3dB

m

m

)

f

/

(f

 

 

1

f

2

 UIpp)

0.05

 

setting

 

of

 

Q%

5

.

2

(

+

p

×

±

×

±


over the range:



10 Hz ( fm ( 2·f3
where fm is the modulation frequency, f3dB = 2·f3 ±10% is the bandwidth of the low-pass filter, f3 is defined in Table 3 for SDH line signals and in Table 3a for SDH tributary signals, and Q is a variable error specified in Table 4 for SDH line signals and in Table 4a for SDH tributary signals. The frequencies f0, f1, f4 and f9 used in Table 4 and Table 4a are defined in Table 3 and Table 3a.

See Annex B for the definition of band-limited peak-to-peak phase slope error.

NOTE – This specification includes modulation harmonics (within the low-pass filter bandwidth) due to distortion, but it excludes any wideband intrinsic jitter/wander components.

	Table 4/O.172 – Variable error (Q) of SDH line jitter/wander generation

	Signal
	Error, Q
	Frequency range

	STM‑0e, STM‑0
	FFS
	f9 ( f1

	
	(8%
	f1 ( f4

	STM‑1e, STM‑1
	FFS
	f0 ( f1

	
	(8%
	f1 ( 500 kHz

	
	(12%
	500 kHz ( f4

	STM‑4, STM‑16, STM‑64
	FFS
	f0 ( f1

	
	(8%
	f1 ( 500 kHz

	
	(12%
	500 kHz ( 2 MHz

	
	(15%
	2 MHz ( f4

	STM-256
	FFS
	FFS

	NOTE – FFS denotes that the value is for further study.


Table 4a/O.172 – Variable error (Q) of SDH tributary jitter/wander generation

	Bit rate (kbit/s)
	Error, Q
	Frequency range

	1544, 2048, 6312, 44 736
	(8%
	f1 ( f4

	34 368
	(8%
	f1 ( 500 kHz

	
	(12%
	500 kHz ( f4

	139 264
	(8%
	f1 ( 500 kHz

	
	(12%
	500 kHz ( 2 MHz

	
	(15%
	2 MHz ( f4


8.6.3
Intrinsic jitter/wander of generation function

The intrinsic jitter of the jitter/wander generation function measured in a bandwidth f1 ( f4 as defined in Table 7 with the amplitude set to zero shall be less than:


0.04 UIpp
for an output signal with structure defined in Annex A; or


0.02 UIpp
for a clock signal.

The specification for maximum allowable intrinsic wander is for further study.

9
Jitter measurement function

9.1
Reference timing signal

A reference timing signal for the phase detector is required. For end-to-end measurements of jitter it may be derived in the jitter measurement function from the input digital test pattern. For looped measurements it may be derived from a suitable clock source.

9.2
Measurement capabilities

9.2.1
Measurement range

The jitter measurement function shall be capable of measuring peak-to-peak jitter. The measurement ranges to be provided are optional but for reasons of compatibility the jitter amplitude/jitter frequency characteristic of the jitter measurement function shall meet the minimum requirements of Figure 4 and Table 5 for SDH line signals or Table 5a for SDH tributary signals. The frequencies f6 to f4 define the range of jitter frequencies to be measured; capability to measure the range of frequencies lower than f1 is optional.

NOTE – Operation of the jitter measurement function over one continuous frequency range f6 to f4 is optional.

Table 5/O.172 – Minimum amplitude of measured jitter versus jitter frequency

	Signal
	Minimum peak-to-peak jitter amplitude (UIpp)
	Jitter frequency breakpoints (Hz)

	
	A2
	A3
	A4
	f6
	f7
	f1
	f2
	f3
	f4

	STM‑0e,
STM‑0
	
20
	2
	0.2
	10
	30
	300
	2 k
	20 k
	400 k

	STM‑1e
	
50
	2
	0.1
	10
	19.3
	500
	3.25 k
	65 k
	1.3 M

	STM‑1
	
50
	2
	0.2
	10
	19.3
	500
	6.5 k
	65 k
	1.3 M

	STM‑4
	
200
	2
	0.2
	*
	10
	1 k
	25 k
	250 k
	5 M

	STM‑16
	
800
	2
	0.2
	10
	12.1
	5 k
	100 k
	1 M
	20 M

	STM‑64
	
3200
	2
	0.2
	10
	12.1
	20 k
	400 k
	4 M
	80 M

	STM-256
	FFS
	FFS
	FFS
	FFS
	FFS
	FFS
	FFS
	FFS
	FFS

	NOTE 1 – Values denoted by "*" are not defined.

NOTE 2 – The accuracy of the instrument is specified between frequencies f1 and f4.

NOTE 3 – Values for STM‑0 are based on the requirements of ANSI T1.105.03 [20].


Table 5a/O.172 – Minimum measured jitter amplitude versus jitter frequency 

	Signal (kbit/s)
	Minimum peak-to-peak jitter amplitude (UIpp)
	Jitter frequency breakpoints (Hz)

	
	A2
	A3
	A4
	f6
	f7
	f1
	f2
	f3
	f4

	1544
	*
	10
	0.5
	*
	*
	10
	400
	8 k
	40 k

	2048
	*
	10
	0.5
	*
	*
	20
	900
	18 k
	100 k

	6312
	*
	10
	0.5
	*
	*
	10
	

1600
	32 k
	60 k

	34 368
	*
	10
	0.5
	*
	*
	100
	

1000
	20 k
	800 k

	44 736
	*
	10
	0.5
	*
	*
	10
	

5000
	100 k
	400 k

	139 264
	*
	10
	0.5
	*
	*
	200
	500
	10 k
	3500 k

	NOTE 1 – Values denoted by "*" are not defined.

NOTE 2 – The accuracy of the instrument is specified between frequencies f1 and f4.



[image: image6.wmf]T0409230-98

A

2

A

3

A

4

f

7

f

1

f

2

f

3

f

4

f

6

Peak-to-peak

jitter/wander amplitude

(log scale)

Frequency (log scale)


Figure 4/O.172 – Measured jitter amplitude versus jitter frequency

9.2.2
Selectable threshold

When measuring peak-to-peak jitter it shall be possible to count the number of occasions and the period of time for which a given selectable threshold of jitter is exceeded. It shall be possible to record these events by means of an external counter, or an internal counter as an option.

It shall be possible to set the threshold at any selected value within the measuring range of the jitter measurement function.

9.2.3
Measurement of RMS jitter

Measurement of RMS jitter may be performed internally within the instrumentation, or externally using the analogue output mentioned in 9.5.1.

9.2.4
Input phase tolerance for SDH tributary signals

The test set shall tolerate input sinusoidal phase variation at particular tributary bit rates, according to the following frequency/amplitude specifications of Table 6. These frequency/amplitude values represent the relevant worst-case pointer test sequences defined in ITU‑T G.783 [5].

Table 6/O.172 – Input phase tolerance of test
set when measuring SDH tributary jitter

	Bit rate
(kbit/s)
	Input phase variation

	
	Amplitude 
(UIpp)
	Frequency 
(Hz)

	1544
	17
	3.0

	2048
	30
	0.5

	6312
	FFS
	FFS

	34 368
	22
	5.0

	44 736
	60
	1.5

	139 264
	75
	1.5

	NOTE – FFS denotes that the value is for further study.


In this context, tolerate means that the test set shall measure within the accuracy specified in this Recommendation, considering that the input phase variation may have been attenuated by the applied measurement filters.

9.3
Measurement bandwidths

The measurement bandwidth shall be limited in order to measure the specified jitter spectra as defined in relevant Recommendations and for other uses. The bandwidth f1 ( f4 or f3 ( f4 of the jitter measurement function shall be in accordance with Table 7 for SDH line signals and Table 7a for SDH tributary signals.

Table 7/O.172 – Jitter measurement function bandwidth for SDH line signals

	Signal
	Jitter measurement bandwidth
((3 dB cut-off frequencies)

	
	f1 (Hz)
high-pass
	f3 (Hz)
high-pass
	f4 (Hz)
low-pass

	STM‑0e, STM‑0
	100
	20 k
	400 k

	STM‑1e, STM‑1
	500
	65 k
	1.3 M

	STM‑4
	1 k
	250 k
	5 M

	STM‑16
	5 k
	1 M
	20 M

	STM‑64
	20 k
	4 M
	80 M

	STM-256
	80 k
	16 M
	320 M

	NOTE 1 – Values for STM‑0 are based on the requirements of ANSI T1.105.03 [20].

NOTE 2 – Values for STM‑256 are to be considered provisional, since network requirements are not yet defined in ITU‑T G.825 [12].


Table 7a/O.172 – Jitter measurement function bandwidth for SDH tributary signals

	Bit rate
(kbit/s)
	Jitter measurement bandwidth
((3 dB cut-off frequencies)

	
	f1 (Hz)
high-pass
	f3 (Hz)
high-pass
	f4 (Hz)
low-pass

	1544
	10
	8 k
	40 k

	2048
	20
	18 k (0.7 k)
	100 k

	6312
	10
	3 k
	60 k

	34 368
	100
	10 k
	800 k

	44 736
	10
	30 k
	400 k

	139 264
	200
	10 k
	3.5 M

	NOTE – Two values are specified for f3 at 2048 kbit/s. The value shown in parenthesis only apply to measurements at certain national interfaces.


9.3.1
Frequency response of jitter measurement function for SDH line signals

The response of all filters within the pass band shall be such that the accuracy requirements of the jitter measurement function are met (refer to 9.4).

For all SDH line bit rates, the following requirements apply to the jitter measurement function when the measurement filters at frequencies f1, f3 and f4 are used:

a)
The high-pass measurement filters with cut-off frequencies f1 and f3 have a first-order characteristic and a roll-off of 20 dB/decade.

b)
The nominal f1 and f3 cut-off frequencies for each bit rate are specified in Table 7 and the nominal –3 dB point of the measurement filters shall be at frequencies f1 (10% and f3 (10%, respectively.

c)
The low-pass measurement filter with cut-off frequency f4 has a maximally-flat, Butterworth characteristic and a roll-off of (60 dB/decade.

d)
The nominal f4 cut-off frequency for each bit rate is specified in Table 7 and the (3 dB point of the measurement filter shall be at a frequency f4 (10%.

e)
The maximum attenuation of the measurement filters shall be at least 60 dB.

These jitter measurement functional requirements are compatible with ITU‑T G.825 [12].

9.3.2
Frequency response of jitter measurement function for SDH tributary signals

The response of all filters within the pass band shall be such that the accuracy requirements of the jitter measurement function are met (refer to 9.4).

For all SDH tributary bit rates, the following requirements apply to the jitter measurement function when the measurement filters at frequencies f1, f3 and f4 are used:

a)
The high-pass measurement filters with cut-off frequency f1 or f3 have a first-order characteristic and a roll-off of 20 dB/decade.

b)
The nominal f1 and f3 cut-off frequency for each bit rate is specified in Table 7a and the nominal (3 dB point of the measurement filters shall be at a frequency f1 (10% and f3 (10% respectively.

c)
The low-pass measurement filter with cut-off frequency f4 has a maximally-flat, Butterworth characteristic and a roll-off of at least (60 dB/decade for bit rates 2048, 34 368 and 139 264 kbit/s and (20 dB/decade for bit rates 1544, 6312 and 44 736 kbit/s.

d)
The nominal f4 cut-off frequency for each bit rate is specified in Table 7a and the nominal (3 dB point of the measurement filter shall be at a frequency f4 ±10%.

e)
The maximum attenuation of the measurement filters shall be at least 60 dB.

f)
In addition, the following requirements apply when the high-pass filter function is at f1. The variable error of the measurement filters below f1 shall be as defined in Table 10a for frequency f1. A second filter pole is allowed at a frequency of less than 0.1 Hz, where the roll-off characteristic can increase to 40 dB/decade.

These jitter measurement functional requirements are compatible with ITU‑T G.783 [5], ITU‑T G.823 [10], and ITU‑T G.824 [11].

Appendix IV illustrates how these filter requirements and the specifications of the following subclauses may be combined into a total jitter measurement function response.

9.4
Measurement accuracy

9.4.1
Measurement result accuracy

The measuring accuracy of the jitter measurement function is dependent upon several factors such as fixed intrinsic error, frequency response and digital test pattern-dependent error of the internal reference timing circuits. In addition, there is an error that is a function of the actual reading.

The accuracy of the jitter measurement shall not be affected by frequency offset on the input signal that is within the limits defined for the various bit rates in ITU‑T G.703 [2], ITU‑T G.813 [9] and ITU‑T G.958 [14].

The measurement accuracy is specified using an input signal with structure defined in Annex A for SDH line signals or 8.3.2 for pseudorandom sequences for SDH tributary signals and physical characteristics of either:

a)
an electrical signal in conformance with ITU‑T G.703 [2] having the nominal terminated signal level and with no additional frequency-dependent loss; or

b)
an optical signal in conformance with ITU‑T G.957 [13] or ITU‑T G.691 [1] and with a nominal power in the range (10 dBm to (12 dBm.

The total measurement error shall be less than:


(R% of reading (W

where R is the variable error specified in Table 10 or Table 10a and W is the fixed error of Table 8 or Table 9, which includes any contribution from the internal timing extraction function.

9.4.2
Fixed error of SDH line jitter measurements

For the STM‑N bit rates and for the indicated digital signals, the fixed error of the jitter measurement function shall be as specified in Table 8 within the frequency ranges f1 ( f4 and f3 ( f4 indicated. Frequencies f1, f3 and f4 used in Table 8 are defined in Table 7.

Table 8/O.172 – Fixed error (W) of SDH line jitter measurements

	Signal
	Maximum peak-to-peak jitter error (UIpp)
for given digital signals

	
	Structured signal
	Clock signal

	
	f1 ( f4
	f3 ( f4
	f1 ( f4
	f3 ( f4

	STM‑0e
	FFS
	FFS
	FFS
	FFS

	STM‑0
	0.07
	0.05
	0.05
	0.03

	STM‑1e
	0.07
	0.025
	0.05
	0.02

	STM‑1
	0.07
	0.05
	0.05
	0.03

	STM‑4
	0.1
	0.05
	0.05
	0.03

	STM‑16
	0.1
	0.05
	0.05
	0.03

	STM‑64
	0.15
	0.05
	0.05
	0.03

	STM-256
	FFS
	FFS
	FFS
	FFS

	NOTE 1 – FFS denotes that the value is for further study.

NOTE 2 – Structured digital signals are defined in Annex A.

NOTE 3 – Clock interfaces are optional.


9.4.3
Fixed error of SDH tributary jitter measurements

For the tributary bit rates and for the indicated digital signals, the fixed error of the jitter measurement function shall be as specified in Table 9 within the frequency ranges f1 ( f4 and f3 ( f4 indicated. Frequencies f1, f3 and f4 used in Table 9 are defined in Table 7a.

Table 9/O.172 – Fixed error (W) of SDH tributary jitter measurements

	Bit rate
(kbit/s)
	Maximum peak-to-peak jitter error (UIpp)
for given digital signals

	
	Pseudo-random signal
	Clock signal

	
	f1 ( f4
	f3 ( f4
	f1 ( f4
	f3 ( f4

	1544
	0.04
	0.025
	0.015
	0.01

	2048
	0.04
	0.025
	0.015
	0.01

	6312
	0.04
	0.025
	0.015
	0.01

	34 368
	0.04
	0.025
	0.03
	0.02

	44 736
	0.04
	0.025
	0.03
	0.02

	139 264
	0.04
	0.025
	0.03
	0.02

	NOTE 1 – Pseudo-random digital signals are defined in 8.3.2.

NOTE 2 – Clock interfaces are optional.


9.4.4
Variable error of SDH tributary jitter measurements 

At jitter frequencies between f1 and f4, the variable error R additional to that specified in 9.3.1 shall be as specified in Table 10 for SDH line signals and Table 10a for SDH tributary signals. Frequencies f1 and f4 used in Table 10 are defined in Table 7. Frequencies f1 and f4 used in Table 10a are defined in Table 7a.

Table 10/O.172 – Variable error (R) of SDH line jitter measurements

	Signal
	Error, R
	Frequency range

	STM‑0e, STM‑0
	FFS
	f1 ( f4

	STM‑1e, STM‑1
	(7%
	f1 ( 300 kHz

	
	(8%
	300 kHz ( 1 MHz

	
	(10%
	1 MHz ( f4

	STM‑4
	(7%
	f1 ( 300 kHz

	
	(8%
	300 kHz ( 1 MHz

	
	(10%
	1 MHz ( 3 MHz

	
	(15%
	3 MHz ( f4

	STM‑16, STM‑64
	(7%
	f1 ( 300 kHz

	
	(8%
	300 kHz ( 1 MHz

	
	(10%
	1 MHz ( 3 MHz

	
	(15%
	3 MHz ( 10 MHz

	
	(20%
	10 MHz ( f4

	STM-256
	FFS
	FFS

	NOTE – FFS denotes that the value is for further study.


Table 10a/O.172 – Variable error (R) of SDH tributary jitter measurements

	Bit rate
(kbit/s)
	Error R
	Frequency range

	1544
	±9%
	f1 – 1 kHz

	
	±7%
	1 kHz – f4

	2048
	±7%
	f1 – f4

	6312
	±9%
	f1 – 1 kHz

	
	±7%
	1 kHz – f4

	34 368
	±7%
	f1 – 300 kHz

	
	±8%
	300 kHz – f4

	44 763
	±9%
	f1 – 200 Hz

	
	±7%
	200 Hz – 300 kHz

	
	±8%
	300 kHz – f4

	139 264
	±7%
	f1 – 300 kHz

	
	±8%
	300 kHz – 1 MHz

	
	±10%
	1 MHz – 3 MHz

	
	±15%
	3 MHz – f4


9.4.5
Digital test signal-dependent error

The accuracy requirements stated in previous subclauses shall be met when digital test signals defined in Annex A are used to perform the jitter measurement. When using other structured signals, pseudo-random or random signals, larger measurement errors could be expected. Considering the measurement bandwidths specified above, signals with a higher "zero" or "one" content (i.e. fewer signal transitions) may even infringe the sampling theorem which, for theoretical reasons, makes it impossible to meet the specified accuracy requirements.

9.5
Additional facilities

9.5.1
Analogue output

The jitter measurement function may provide an analogue output signal to enable measurements to be made externally to the jitter measurement function, e.g. by using an oscilloscope or an RMS meter.

10
Wander measurement function

Appendix II provides further information regarding the test configurations for measurement of wander.

10.1
Reference timing signal

For the testing of wander, it shall be possible to phase-lock the measurement function to an external reference clock source of arbitrary accuracy; refer also to 7.3.

10.2
Measurement of TIE (Time Interval Error)

The instrumentation shall be capable of measuring Time Interval Error (TIE) as defined in ITU‑T G.810 [6]. TIE is the basic function whereby many different stability parameters (such as MTIE and TDEV) may be calculated.

TIE can be interpreted as the time difference between the signal being measured and the reference clock. It is typically measured in nanoseconds and set to zero at the start of the measurement period. Therefore, TIE gives the timing change since the measurement began.

10.2.1
Sampling interval

In order to calculate and estimate the various wander parameters specified in the following subclauses, TIE is treated as a sampled parameter since continuous knowledge of the time interval error is not practically attainable (refer to ITU‑T G.810 [6]).

The maximum sampling time (0, of TIE shall be:


1/30 s

in accordance with ITU‑T G.813 [9] and ITU‑T G.812 [8].

10.2.2
Measurement bandwidth

Wander shall be measured through an equivalent 10 Hz, first-order, low-pass measurement filter, in accordance with ITU‑T G.813 [9] and ITU‑T G.812 [8], and with the following characteristics:

a)
The low-pass measurement filter has a single-order characteristic and a roll-off of  EQ  (20 dB/decade. The (3 dB point of the measurement filter shall be at a frequency 10 Hz (10%.

b)
The amplitude of pass-band ripple in the range 1 to 10 Hz shall be less than (0.2 dB (relative to the gain at 0.1 Hz) and the maximum attenuation of the measurement filter shall be at least 30 dB.

10.2.3
Measurement range

The dynamic range of the TIE measurement shall be a minimum of:


(1 ( 109 ns

(corresponding, for example, to a frequency offset of (4.6 ppm for 200 000 seconds or over 55 hours).

10.2.4
Measurement result accuracy

The measuring accuracy of the wander measurement function is dependent upon several factors such as the magnitude of the reading, fixed intrinsic error, frequency response and TIE sampling interval.

For each measurement of TIE over an Observation Interval (, the total TIE measurement error shall be less than:


(5% of the measured TIE value (Z0(()

where Z0(() is based on the measurement requirements of ITU‑T G.811 [7] and is specified in Table 11.

Table 11/O.172 – Fixed error (Z0) of TIE measurement

	Z0(() (ns)
	Observation Interval, ( (s)

	2.5 ( 0.0275 (
	0.05 ( ( ( 1000

	29 ( 0.001 (
	( ( 1000

	NOTE – There is an additional frequency-dependent error term above 1 Hz due to the measurement filter response (refer to 10.2.2).


10.3
Measurement of transient TIE (Time Interval Error)

The instrument may be capable of measuring transient Time Interval Error (TIE). Transient TIE is defined as TIE measured through an equivalent 100 Hz, first-order, low-pass measurement filter as described in ITU‑T G.783 [5] and ITU‑T G.813 [9].

Transient TIE may be used for measurement of pointer adjustments specified in ITU‑T G.783 [5] or for clock phase transients specified in ITU‑T G.812 [8] and ITU‑T G.813 [9].

10.3.1
Sampling Interval

The maximum sampling time (0, of transient TIE shall be:


1/1000 s

in accordance with ITU‑T G.783 [5] and ITU‑T G.812 [8].

10.3.2
Measurement Bandwidth

Transient TIE shall be measured through an equivalent 100 Hz, first-order, low-pass measurement filter.

a)
The low-pass measurement filter has a single-order characteristic and a roll-off of  EQ  –20 dB/decade. The (3 dB point of the measurement filter shall be at a frequency 100 Hz (10%.

b)
The amplitude of pass-band ripple in the range 10 to 100 Hz shall be less than (0.2 dB (relative to the gain at 1 Hz) and the maximum attenuation of the measurement filter shall be at least 30 dB.

10.3.3
Measurement Range

The dynamic range of the transient TIE measurement shall be a minimum of:


(1 ( 106 ns

10.3.4
Measurement Result Accuracy

The measuring accuracy of the transient TIE measurement function is dependent upon several factors such as the magnitude of the reading, fixed intrinsic error, frequency response and sampling interval.

For each measurement of transient TIE over an Observation Interval (, the total measurement error shall be less than:


(5% of the measured transient TIE value (Z9(()

where Z9(() is based on the measurement requirements of ITU‑T G.811 [7] and is specified in Table 12.

Table 12/O.172 – Fixed error (Z9) of transient TIE measurement

	Z9(() (ns)
	Observation Interval, ( (s)

	2.5 ( 0.0275 (
	0.001( ( ( 100

	NOTE – There is an additional frequency-dependent error term above 10 Hz due to the measurement filter response (refer to 10.3.2).


10.4
Measurement of MTIE (Maximum Time Interval Error)

The capability of measuring Maximum Time Interval Error (MTIE) as defined in ITU‑T G.810 [6] may be provided.

MTIE is a measure of wander that characterizes frequency offsets and phase transients. It is a function of a parameter ( called the Observation Interval. MTIE(() can be said to be the largest peak‑to-peak TIE in any observation interval of length (.

10.4.1
Measurement and observation interval ranges

In order to support the MTIE specifications of various ITU‑T Recommendations, it shall be possible to measure MTIE over a range of observation intervals at least from:


0.05 s to 100 000 s for TIE as described in 10.2;


0.001 s to 100 s for transient TIE as described in 10.3.

The maximum range of calculated MTIE results shall be at least:


50 000 ns

(corresponding to 10 times the maximum specification defined in ITU‑T G.813 [9], for example).

NOTE – The minimum measurement period T for MTIE(() is the observation interval (i.e. T = ().

10.4.2
Calculation algorithm accuracy

In certain cases, the MTIE calculation algorithm can be functionally separated from the TIE measurement, in which case the following accuracy requirements apply to the standalone algorithm.

When provided with a given set of TIE measurement data, an algorithm used to calculate MTIE shall yield results within a certain error of the values calculated in accordance with the standard estimator formula given in II.5/G.810 [6].

The total MTIE calculation error shall be less than:


(2% of the MTIE value (Z1(()

where Z1(() is specified in Table 13 and ( is the observation interval.

Table 13/O.172 – Fixed error (Z1) of MTIE calculation algorithm

	Z1(() (ns)
	Observation Interval, ( (s)

	0.5 ( 0.0055 (
	0.001 ( ( ( 1000

	5.8 ( 0.0002 (
	( ( 1000

	NOTE – These requirements are based on a consideration of ITU‑T G.811 [7].


In order to verify the accuracy of a standalone MTIE calculation algorithm, a defined TIE noise source can be used, which is described in Appendix V.

10.4.3
Measurement result accuracy

The total measurement error (i.e. including error from a TIE measurement and error from the MTIE calculation algorithm) shall be less than:


(7% of the MTIE value (Z3(()

where Z3(() is specified in Table 14 and ( is the observation interval.

Table 14/O.172 – Fixed error (Z3) of MTIE measurement result

	Z3(() (ns)
	Observation Interval, ( (s)

	3 ( 0.033 (
	0.001 ( ( ( 1000

	35 ( 0.0012 (
	( ( 1000

	NOTE – These requirements are based on a consideration of ITU‑T G.811 [7].


10.5
Measurement of TDEV (Time Deviation)

The capability of measuring Time Deviation (TDEV) as defined in ITU‑T G.810 [6] may be provided.

TDEV is a measure of wander that characterizes its spectral content. It is a function of a parameter ( called the Observation Interval. TDEV(() can be said to be the RMS of filtered TIE, where a band‑pass filter is centred on a frequency of 0.42/(.

10.5.1
Measurement and observation interval ranges

In order to support the TDEV specifications of various ITU‑T Recommendations, it shall be possible to measure TDEV over a range of observation intervals at least from:


0.05 s to 10 000 s

The maximum range of calculated TDEV results shall be at least:


10 000 ns

(corresponding to 10 times the maximum specification defined in ITU‑T G.813 [9], for example).

NOTE – The minimum measurement period T for TDEV(() is twelve times the observation interval (i.e. T = 12(), in accordance with ITU‑T G.813 [9], ITU‑T G.812 [8] and ITU‑T G.811 [7].

10.5.2
Calculation algorithm accuracy

In certain cases, the TDEV calculation algorithm can be functionally separated from the TIE measurement, in which case the following accuracy requirements apply to the standalone algorithm.

When provided with a given set of TIE measurement data, an algorithm used to calculate TDEV shall yield results within a certain error of the values calculated in accordance with the standard estimator formula given in II.3/G.810 [6].

The total TDEV calculation error shall be less than:


(2% of the TDEV value (Z2(()

where Z2(() is specified in Table 15 and ( is the observation interval.

Table 15/O.172 – Fixed error (Z2) of TDEV calculation algorithm

	Z2(() (ns)
	Observation Interval, ( (s)

	0.06
	0.05 ( ( ( 100

	0.0006 (
	100 ( ( ( 1000

	0.6
	1000 ( ( ( 10 000

	NOTE – These requirements are based on a consideration of ITU‑T G.811 [7].


In order to verify the accuracy of a standalone TDEV calculation algorithm, a defined TIE noise source can be used, which is described in Appendix V.

10.5.3
Measurement result accuracy

The total measurement error (i.e. including error from a TIE measurement and error from the TDEV calculation algorithm) shall be less than:


(7% of the TDEV value (Z4(()

where Z4(() is specified in Table 16 and ( is the observation interval.

Table 16/O.172 – Fixed error (Z4) of TDEV measurement result

	Z4(() (ns)
	Observation Interval, ( (s)

	0.36
	0.05 ( ( ( 100

	0.0036 (
	100 ( ( ( 1000

	3.6
	1000 ( ( ( 10 000

	NOTE – These requirements are based on a consideration of ITU‑T G.811 [7].


10.6
Measurement of Frequency Offset 

The capability of measuring frequency offset, as defined below, may be provided.

Frequency offset is the first time derivative of phase. The transfer function of the derivative process grows in proportion to the frequency with which the frequency offset varies. Therefore the bandwidth of the measurement must be limited. The measurement bandwidth is controlled by the time period over which the frequency offset is calculated from a set of TIE values xi.

Fractional frequency offset (in ns/s) measured over a period T is defined as




[image: image7.wmf]å

=

+

ú

û

ù

ê

ë

é

-

-

-

t

=

t

N

i

i

n

N

N

i

x

N

n

y

1

2

0

0

1

1

1

2

6

)

(

,

where:


(0
is the sampling interval in seconds,


N
is the number of phase samples in the measurement period,


T = N(0
is the measurement period in seconds,


t = n(0
is the time at the beginning of the measurement period,


xi
are the phase samples in ns.

NOTE – This measurement attenuates frequency offset components that vary at frequencies greater than 0.55/T. The user must choose the value of T = N(0 to attenuate noise and pass the frequencies of interest.

10.6.1
Measurement range

In order to support the frequency offset measurement specifications of various ITU‑T Recommendations, it shall be possible to measure frequency offset over a measurement period range of at least:


0.05 s ( T ( 10 000 s.

The maximum range of calculated frequency offset results shall be at least:


75 000 ns/s

(corresponding to 10 times the maximum specification in ITU‑T G.813 [9]).

10.6.2
Calculation algorithm accuracy

In certain cases, the frequency offset calculation algorithm can be functionally separated from the TIE measurement, in which case the following accuracy requirements apply to the standalone algorithm.

When provided with a given set of TIE measurement data, an algorithm used to calculate frequency offset shall yield results within a certain error of the values calculated in accordance with the standard formula given in 10.6.

The total frequency offset calculation error shall be less than:


(2% of the frequency offset value (Z5(T)

where Z5(T) is specified in Table 17 and T is the measurement period.

Table 17/O.172 – Fixed error (Z5) of frequency offset calculation algorithm

	Z5(T) (ns/s)
	Measurement Period, T(sec)

	0.0055
	0.05 ( T ( 1000

	0.0002
	T ( 1000

	NOTE – These requirements are based on a consideration of ITU‑T G.811 [7].


10.6.3
Measurement result accuracy

The total measurement error (i.e. including error from a TIE measurement and error from the frequency offset calculation algorithm) shall be less than:


±7 % of the frequency offset value ±Z6(T)

where Z6(T) is specified in Table 18 and T is the measurement period.

Table 18/O.172 – Fixed error (Z6) of frequency offset measurement result

	Z6(T) (ns/s)
	Measurement Period, T(sec)

	0.033
	0.05 ( T ( 1000

	0.0012
	T ( 1000

	NOTE – These requirements are based on a consideration of ITU‑T G.811 [7].


10.7
Measurement of Frequency Drift Rate 

The capability of measuring frequency drift rate, as defined below, may be provided.

Frequency drift rate is the second time derivative of phase. The transfer function of the second-derivative process grows in proportion to the square of the frequency with which the frequency drift rate varies. Therefore the bandwidth of the measurement must be limited. The measurement bandwidth is controlled by the time period over which the frequency drift rate is calculated from a set of TIE values xi.

Frequency drift rate (in ns/s2) measured over a period T is defined as




[image: image8.wmf]å

=

+

ú

ú

û

ù

ê

ê

ë

é

+

-

+

+

-

-

-

+

-

t

=

t

N

i

i

n

N

N

N

N

N

i

N

N

i

x

N

n

D

1

2

2

3

2

4

2

2

0

0

2

3

1

4

4

6

4

5

6

60

)

(


where:


(0
is the sampling interval in seconds,


N
is the number of phase samples in the measurement period,


T = N(0
is the measurement period in seconds,


t = n(0
is the time at the beginning of the measurement period,


xi
are the phase samples in ns.

NOTE – This measurement attenuates frequency drift rate components that vary at frequencies greater than 0.8/T. The user must choose the value of T = N(0 to attenuate noise and pass the frequencies of interest.

10.7.1
Measurement range

In order to support the frequency drift rate measurement specifications of various ITU‑T Recommendations, it shall be possible to measure frequency drift rate over a measurement period range of at least:


0.05 s ( T ( 10 000 s.

The maximum range of calculated frequency drift rate results shall be at least:


0.06 ns/s2

(corresponding to 10 times the maximum specification in ITU‑T G.813 [9]).

10.7.2
Calculation algorithm accuracy

In certain cases, the frequency drift rate calculation algorithm can be functionally separated from the TIE measurement, in which case the following accuracy requirements apply to the standalone algorithm.

When provided with a given set of TIE measurement data, an algorithm used to calculate frequency drift rate shall yield results within a certain error of the values calculated in accordance with the standard formula given in 10.7.

The total frequency drift rate calculation error shall be less than:


(2% of the frequency drift rate value (Z7(T)

where Z7(T) is specified in Table 19 and T is the measurement period.

Table 19/O.172 – Fixed error (Z7) of frequency drift rate calculation algorithm

	Z7(T) (ns/s2)
	Measurement Period, T (sec)

	0.5·T–2
	0.05 < T < 2500

	8·10–8
	T > 2500

	NOTE – These requirements are based on a consideration of ITU‑T G.812 [8].


10.7.3
Measurement result accuracy

The total measurement error (i.e. including error from a TIE measurement and error from the frequency drift rate calculation algorithm) shall be less than:


±7 % of the frequency drift rate value ±Z8(T)

where Z8(T) is specified in Table 20 and T is the measurement period.

Table 20/O.172 – Fixed error (Z8) of frequency drift rate measurement result

	Z8(T) (ns/s2)
	Measurement Period, T(sec)

	T–2
	0.05 < T < 2500

	1.6·10–7
	T > 2500

	NOTE – These requirements are based on a consideration of ITU‑T G.812 [8].


11
TDEV wander noise generation function

The capability of generating TDEV wander noise for wander tolerance and wander transfer measurements as described in ITU-T G.812 [8] and ITU‑T G.813 [9] may be provided. To ensure sufficiently accurate, robust and consistent measurements, the following requirements shall be met:

a)
The TDEV noise generator shall produce a test signal within ±20% of the applicable TDEV input noise tolerance mask.

b)
The skewness S and the kurtosis K of generated TDEV wander noise must satisfy 



–0.2 ( S(τ,T) ( 0.2


and 


2.5 ( K(τ,T) ( 3.5


for τ ( 10 000 seconds and T ( 12τ, where τ is the integration period and T is the measurement period. These parameters are defined in Annex C.

c)
The test signal shall be deterministic and repeatable, which implies that the signal is able to start at the same point of the signal time function.

Appendix II.2/G.812 [8] shows an example for an adequate implementation algorithm for generating a TDEV wander noise signal.

Appendix II.1/G.812 [8] has additional information on the application of the TDEV wander noise signal.

NOTE – The signal may exceed the corresponding MTIE tolerance mask limits. Since both requirements TDEV and MTIE have to be met for an input interface, the TDEV limits may be more stringent and therefore exceeding the MTIE values may be tolerable. The maximum allowance for exceeding the MTIE limits is for further study.

12
Operating environment

The performance requirements shall be met when operating within the climatic conditions as specified in 2.1/O.3 [15].

Annex A

Structured test signals for the measurement of jitter

A.1
Introduction

It is important to define the test signals to be used when performing jitter tests. This is of particular importance when testing SDH optical systems, since the scrambling system does not limit the length of runs of zeros/ones that can exist on a line signal, i.e. the maximum time period without any data transitions in the scrambled signal. ITU‑T G.707 [3] provides further information regarding SDH signal structure and payload scrambling.

For example, if the traffic in an STM‑N signal emulates the scrambling pattern, then many bytes of all zeros/ones will appear in the coded line signal. As the extreme cases of this will be rare occurrences, and also as it will be very difficult for a jitter test set to continue to perform accurate measurements under these conditions, it is important that a representative worst-case signal is defined for the purposes of test set specification.

A.1.1
Payload test conditions

Concatenated payloads provide the worst-case scenario for STM‑N test signals. For bulk-filled concatenated signals with a 223 ( 1 PRBS filling the container, the result of scrambling this data is a worst-case run of 30 consecutive identical zeros/ones (i.e. there will be 30 clock periods with no transitions on the line signal). For non-concatenated payloads generated by SDH test sets, the byte interleaving of the VC‑4 containers reduces the maximum length of runs produced.

A.1.2
SDH overhead byte conditions

Care shall be taken in selecting the binary content of the J0 and Z0 bytes and of the bytes reserved for national use which are excluded from the scrambling process of the STM‑N signal, to ensure that long sequences of "1"s or "0"s do not occur.

The content of these overhead bytes shall be set to a "10101010" (AAH) pattern.

A.2
Test signal structure for STM‑N signals

A.2.1
STM‑1 signal

The STM‑1 test signal structure illustrated in Figure A.1 consists of a PRBS test sequence of length 223 ( 1 bits according to ITU‑T O.150 [16], which is applied to all bytes of the C‑4 container.

NOTE – This is equivalent to Test Signal Structure 1 (TSS1) defined in Annex C/O.181 [18].


[image: image9.wmf]T0409240-98

RSOH

AU-PTRs

MSOH

POH

PRBS test

sequence

C-4


Figure A.1/O.172 – Test signal structure
for jitter testing of STM‑1 interface

A.2.2
STM‑N signal (N ( 4)

The test signal TSS1 described in A.2.1 shall be used.

The (N‑1) C‑4 payload containers which do not contain the test signal shall contain an all "0" or all "1" fixed byte pattern or may contain an unequipped VC‑4 as defined in ITU‑T G.783 [5].

A.2.3
STM‑0 signal

The STM‑0 test signal structure illustrated in Figure A.2 consists of a PRBS test sequence of length 223 ( 1 bits according to ITU‑T O.150 [16], which is applied to all bytes of the C‑3 container.


[image: image10.wmf]T0409250-98

RSOH

AU-PTRs

MSOH

POH

PRBS test

sequence

C-3


Figure A.2/O.172 – Test signal structure for
jitter testing of STM‑0 interface

A.3
Test signal structure for concatenated STM‑N signals

A.3.1
STM‑N signal (N ( 4)

The STM‑N test signal structure illustrated in Figure A.3 is a PRBS test sequence of length 
223 ( 1 bits according to ITU‑T O.150 [16], which is applied to all bytes of the C‑4‑nc concatenated container.


[image: image11.wmf]T0409260-98

RSOH

AU-PTRs

MSOH

POH

PRBS test

sequence

C-4-nc

(n-1)

Fixed stuff


Figure A.3/O.172 – Test signal structure for jitter
testing of concatenated STM‑N interface

Annex B

Definition of band-limited peak-to-peak phase slope error

When sinusoidal phase modulation is used to stress a phase-locked loop (PLL), the stress is proportional to the peak-to-peak phase when the modulation frequency is greater than the PLL bandwidth, and the stress is proportional to the peak-to-peak phase slope when the modulation frequency is less than the bandwidth. Distortion in the generated sinusoidal modulation can cause additional error of the phase slope. Clause 8.6.2 specifies limits for the phase slope error.

The phase slope is measured by a procedure functionally equivalent to that illustrated in Figure B.1. The differentiator finds the slope of the phase, and the low-pass filter is first-order with a cut-off frequency f3dB, where f3dB = 2·f3 ±10%, and f3 is defined in Table 3 for SDH line signals and in Table 3a for SDH tributary signals.


[image: image12.wmf]T0415160-01

Differentiator

Signal 

under test

Phase

demodulator

First

-

order

low

-

pass 

filter

Peak

-

to

-

peak

detector

Measured

band

-

limited

peak

-

to

-

peak

phase slope


Figure B.1/O.172 – Phase slope measurement function

If the jitter/wander generation were a perfect sinusoid with peak-to-peak phase amplitude (in UIpp) exactly equal to the "setting" and a modulation frequency of fm, then the peak-to-peak phase slope as measured through a first-order low-pass filter would be:



ideal band-limited peak-to-peak phase slope = 
[image: image13.wmf]2

3dB

m

m

)

f

/

(f

 

 

1

f

2

(setting)

+

p

×


Because the actual phase amplitude is not, in general, equal to the phase amplitude setting, the actual modulation frequency is not equal to fm, and the modulation is not a perfect sinusoid, the measured band-limited peak-to-peak phase slope will not equal the ideal band-limited peak-to-peak phase 

slope given above. The difference between the measured band-limited peak-to-peak phase slope and the ideal band-limited peak-to-peak phase slope is defined as the band-limited peak-to-peak phase slope error.

Annex C

Specification of distribution for TDEV wander noise generation

Let xi be a noise-like sequence (sampled at intervals τ0). Let an integrated sequence yi be defined as a function of integration interval τ = n ·τ0:




[image: image14.wmf]å

=

+

=

t

n

j

j

i

i

x

y

1

)

(


The rth central moment of yi(τ) for a measurement period T = N ·τ0 is defined as




[image: image15.wmf](

)

(

)

[

]

å

+

m

-

t

=

t

N

i

r

i

r

y

N

T

m

1

1

,


where the mean is




[image: image16.wmf](

)

å

+

t

=

m

N

i

i

y

N

1

1


The skewness S and kurtosis K as a function of integration interval τ and measurement period T are defined as




[image: image17.wmf](

)

(

)

(

)

(

)

(

)

(

)

T

m

T

m

T

K

T

m

T

m

T

S

,

,

,

,

,

,

,

2

2

4

5

.

1

2

3

t

t

=

t

t

t

=

t


A sequence xi meets the requirements for TDEV wander noise generation if 



–0.2 ( S(τ,T) ( 0.2

and 

2.5 ( K(τ,T) ( 3.5 

for τ ( 10 000 seconds and T ( 12τ. Note that a Gaussian distribution has S = 0 and K = 3.

APPENDIX I

Guidelines concerning the measurement of jitter in SDH systems

Appendix I/O.171 [17], "Guidelines concerning the measurement of jitter", although appropriate for PDH systems, may also be consulted for guidance on the general principles of measuring jitter in SDH systems.

APPENDIX II

Guidelines concerning the measurement of wander in SDH systems

II.1
Wander measurements

II.1.1
General considerations on wander measurement configurations

Due to the low frequency of the phase variations to be evaluated (refer to definition in clause 3), wander is a quantity which requires a special test configuration. When performing jitter measurements, the required reference timing signal is normally produced locally by means of a phase-locked loop (PLL) within the test set; it is derived from the average phase of the signal to be measured. Such a PLL cannot be realized to cope with the requirements of wander measurements.

Therefore, wander measurements always require an external reference clock signal of adequate stability.

Clauses II.1.2 and II.1.3 contain information on test configurations for wander measurements that are in accordance with ITU‑T G.810 [6].

II.1.2
Synchronized wander measurements

Figure II.1 shows in a very general form the functional block diagram required for the synchronized measurement of wander.


[image: image18.wmf]T0409270-98

Reference

clock

Network or

equipment

under test

Wander

measurement

function

Timing reference

Input

port

Output

port


Figure II.1/O.172 – Synchronized wander
measurement configuration

This configuration is applicable if the timing signals required to perform the measurement can be derived from a common reference clock. This means that only loop measurements, where input and output ports of the unit-under-test are accessible at the same location, can be carried out in this way. In this set-up, the measurement result is not affected by phase variations of the reference clock. Thus, the requirements on the stability of the reference clock are not very high and are achievable in portable test instrumentation.

II.1.3
Non-synchronized wander measurements

The block diagram for non-synchronized wander measurements is illustrated in Figure II.2.


[image: image19.wmf]T0409280-98

Reference

clock 1

Input

port

Output

port

Network or

equipment

under test

Wander

measurement

function

Timing reference

Reference

clock 2


Figure II.2/O.172 – Non-synchronized wander
measurement configuration

This configuration is applicable to wander measurements in cases where both input and output ports of the network or equipment under test are not available at the same location (e.g. end-to-end measurements). In this set-up, the measurement result is affected by any frequency/phase drift of the two clocks involved in the measurement. This means that the stability of the two clocks has to be at least one order of magnitude better than the quantity to be measured. Such reference clocks may not be provided in portable test instrumentation in which case synchronization to an external reference is required.

II.2
Clock stability measurements

If the stability of a clock is to be measured, the measurement set-up is similar to that described above in II.1.3. It is illustrated in Figure II.3.


[image: image20.wmf]T0409290-98

Reference

clock

Timing reference

Wander

measurement

function

Output

port

Clock

equipment

under test


Figure II.3/O.172 – Clock stability
measurement configuration

Also in this configuration, the measurement result is affected by any frequency or phase drift of the reference clock involved in the measurement. The same considerations as described in II.1.3 apply in this case.

APPENDIX III

Guidelines concerning the generation of pointer test sequences

The following functional block diagram of Figure III.1 and text, outline a method for generating the pointer test sequences defined in ITU‑T G.783 [5] and which are described in 8.4.1. This figure does not describe a specific implementation.


[image: image21.wmf]T0409300-98

PRBS*

C-12

VC-12

PRBS*

TU-12

C-4

VC-4

VC-4

AU-4

STM-N

STM-N

Reference

clock

SDH NE

under

test

Fixed

mapping

No AU-4

pointer

adjustments

Fixed

mapping

G.783-defined

pointer sequences

Asynchronous

bit-mapping

Lower-order

path example

Higher-order

path example

Asynchronous

bit-mapping

G.783-defined

pointer sequences

SDH

tributary

(PDH)

Jitter

measurement

* PRBS is used as an example;

   other signals are possible

SDH/PDH test set


Figure III.1/O.172 – Functional block diagram
of pointer test sequence generation

In order to generate a ITU-T G.783 [5] pointer test sequence, the following considerations shall be taken into account by the test set capabilities, the test configuration and the user of the test set:

a)
The SDH/PDH test set and the network element should be synchronized from the same reference clock, to eliminate any uncontrolled pointer justification events that would interfere with the results.

b)
To measure combined pointer and mapping jitter, it should be possible to set the frequency of the PDH PRBS within the PDH offset ranges defined in ITU‑T G.703 [2].

c)
To measure only mapping jitter, it should be possible to suppress pointer justification actions.

d)
To generate the bit-stuffing sequence for asynchronous bit-mapping, justifications of a single polarity should be generated at regular intervals. The mapping process shall be in accordance to 15.2.3.1/G.783[5].

e)
The generation of pointer justification events should be independent of the mapping process.

f)
To create lower-order TU‑m pointer test sequences, the higher-order AU‑n pointer should be fixed.

g)
The "pointer sequences" and the "bit-mapping" functions imply frequency offsets between the clocks which drive the various functional blocks.

APPENDIX IV

Total jitter measurement function response

IV.1
Introduction

Clause 9.3.2 specifies filters for SDH tributary jitter measurement. These specifications are in terms of limits on individual filter parameters such as (3 dB cut-off frequency, gain tolerance, pass-band ripple, etc.

These parameters imply that some equivalent limits on the total frequency responses of the jitter measurement functions exist. This appendix illustrates how the individual filter parameters can be combined into a single frequency response, whose characteristic is bounded by mask limits.

IV.2
Measurement filter parameters

Table IV.1 summarizes the measurement filter parameters and their definitions. In the following descriptions of mask limits, if the value of a particular parameter is not specified in a filter description, then the default value given in Table IV.1 should be used.

Table IV.1/O.172 – Summary of measurement filter specification parameters

	Measurement filter parameter
	Parameter definition

	n
	Order of the filter (n ( 1 is first-order, n ( 3 is third-order).

	fx
	Nominal (3 dB cut-off frequency or bandwidth.

	a
	Fractional cut-off frequency tolerance ((). Default: no specification.

	fp
	Second filter pole frequency. Default: 0 for high-pass, ( for low‑pass.

	g
	Fractional gain tolerance of 9.3.2 c).

	r
	Pass-band ripple (() in dB. Default: 0.

	c
	Minimum maximum attenuation in dB. Default: no specification.


IV.3
Mask limits for high-pass measurement filter response

The upper and lower mask limits for the frequency response are Upper(f) and Lower(f) as defined below. These masks apply to the frequency range f ( 10 fx.




[image: image22.wmf](

)

(

)

ï

î

ï

í

ì

-

+

=

,

,

,

)

(

f

U

c

r

f

U

f

Upper



[image: image23.wmf](

)

otherwise

c

f

U

f

f

x

-

<

>




where:




[image: image24.wmf](

)

ú

ú

û

ù

ê

ê

ë

é

-

+

+

×

=

n

x

n

n

f

a

f

f

g

f

U

2

2

)

1

(

)

1

(

log

20

)

(


[IV-1]




[image: image25.wmf](

)

(

)

ï

î

ï

í

ì

¥

-

-

=

,

,

,

)

(

f

L

r

f

L

f

Lower



[image: image26.wmf](

)

otherwise

c

f

L

f

f

x

-

<

>




where:




[image: image27.wmf](

)

ú

ú

û

ù

ê

ê

ë

é

+

×

+

+

-

×

=

2

2

2

2

)

1

(

)

1

(

log

20

)

(

p

n

x

n

n

f

f

f

f

a

f

f

g

f

L


[IV-2]

IV.3.1
SDH tributary jitter measurement high-pass filter

As an example, the SDH tributary jitter measurement filter specified for 2048 kbit/s has parameters n = 1, fx = 20 Hz, a = 0.1, fp = 0.1 Hz, g = 0.07, r = 0, and c = 60 dB. Then the upper and lower mask limits for the jitter measurement function frequency response are as illustrated in Figure IV.1.


[image: image28.wmf]T0415170-01

10

0

-

10

-2

0

-3

0

-4

0

-5

0

-6

0

-7

0

f

x

 = 20 Hz

10

f

x

Gain

(dB)

Frequency (Hz)

0.01

0.1

1

10

100

1000

f

p

 = 0.1 Hz


Figure IV.1/O.172 – Mask limits for 2048 kbit/s
jitter measurement high-pass filter at 20 Hz

In general, when the jitter measurement function response of a test set lies on the Upper or Lower mask, the error of the measurement (peak-to-peak or r.m.s.) is at its maximum magnitude; other measurement function responses will yield less error. The superposition of g = –0.07 with a = +0.10 or to g = +0.07 with a = –0.10 leads to an extension of the cross-over points of the mask limits with the –3 dB level. Therefore the mask limits for the –3 dB level may exceed the a = ±10% limits specified for the nominal –3 dB cut-off frequency.

APPENDIX V

Verification of MTIE and TDEV calculation algorithms

V.1
TIE noise source functional description

In order to verify the accuracy of standalone MTIE and TDEV calculation algorithms (refer to 10.3.2 and 10.4.2), a defined TIE noise source can be used.

The MTIE and TDEV calculation algorithm accuracy specifications (refer to 10.3.2 and 10.4.2) nominally apply to all possible TIE waveforms from which MTIE and TDEV are calculated. Since it is impractical to test for many waveforms, a simple TIE pattern generator can be used for the evaluation. This pattern should have a 1/f2 power spectral density and a probability distribution function that is approximately Gaussian (satisfies the criteria of Annex C).

V.2
First example of TIE noise generator

One TIE generator that meets these criteria is shown in Figure V.1. It is based on a pseudo-random binary sequence (PRBS) generator that is 31 stages long. The output of the PRBS generator is a pseudo-random sequence of 1 s and 0 s. The output of the TIE generator is a series of numbers xi at a rate fs, where the numbers are given some weight in nanoseconds (ns) by the algorithm under test, and fs is the reciprocal of the sampling interval expected by the algorithm (not less than 30 Hz). 

Note that although the system clock is 6.4 kHz, the numbers xi are generated at a lower rate of fs through sub-sampling. The pattern of this TIE noise generator repeats after 671 000 seconds. The numbers zi before scaling have a dynamic range of 99 123 peak-to-peak. The Scaling factor in Figure V.1 is chosen to match the LSB weighting of the algorithm under test so the dynamic range of the output xi corresponds to 50 000 ns peak-to-peak.


[image: image29.wmf]T0415180-01

17

 

¸

2

z

i

x

i

XOR

Accumulator

Sample

Scaling

PRBS generator

(taps at 10, 13, 24 

and 31)

Map

1 

®

 

 

1 

0

 

®

 

–

1

6.4 kHz

clock

Sampling

clock

(f

s

 

³ 

30 Hz)


Figure V.1/O.172 – TIE noise source used for verifying
MTIE and TDEV calculation algorithms

V.3
Second example of TIE noise generator

An alternative TIE generator for testing MTIE and TDEV calculation algorithms is shown in Figure V.2. It differs from that in Figure V.1 principally in the Mapping function, where the w(n) are floating-point numbers from –2.884 to 2.884 defined by:




[image: image30.wmf]ò

¥

-

-

p

=

+

)

(

 

5

.

0

2

2

1

256

5

.

0

n

w

t

dt

e

n


The output pattern xi has a repetition period of 335 500 seconds, and the dynamic range of zi before scaling is 64 167 peak-to-peak. The Scaling factor in Figure V.2 is chosen to match the LSB weighting of the algorithm under test so the dynamic range of the output xi corresponds to 50 000 ns peak-to-peak.


[image: image31.wmf]T0415190-01

8

 

¸8

z

i

x

i

.

.

.

800 Hz

Accumulator

Sample

Scaling

PRBS generator

(taps at 10, 11, 24 

and 31)

6.4 kHz

clock

Sampling

clock

(f

s

 

³ 

30 Hz)

Serial

-

to

-

Parallel 

Converter

Map

0

 

®

  

w(0)

255 

®

w(255)


Figure V.2/O.172 – Alternative TIE noise source used for verifying
MTIE and TDEV calculation algorithms
	SERIES OF ITU-T RECOMMENDATIONS

	Series A
	Organization of the work of ITU-T

	Series B
	Means of expression: definitions, symbols, classification

	Series C
	General telecommunication statistics

	Series D
	General tariff principles

	Series E
	Overall network operation, telephone service, service operation and human factors

	Series F
	Non-telephone telecommunication services

	Series G
	Transmission systems and media, digital systems and networks

	Series H
	Audiovisual and multimedia systems

	Series I
	Integrated services digital network

	Series J
	Cable networks and transmission of television, sound programme and other multimedia signals

	Series K
	Protection against interference

	Series L
	Construction, installation and protection of cables and other elements of outside plant

	Series M
	TMN and network maintenance: international transmission systems, telephone circuits, telegraphy, facsimile and leased circuits

	Series N
	Maintenance: international sound programme and television transmission circuits

	Series O
	Specifications of measuring equipment

	Series P
	Telephone transmission quality, telephone installations, local line networks

	Series Q
	Switching and signalling

	Series R
	Telegraph transmission

	Series S
	Telegraph services terminal equipment

	Series T
	Terminals for telematic services

	Series U
	Telegraph switching

	Series V
	Data communication over the telephone network

	Series X
	Data networks and open system communications

	Series Y
	Global information infrastructure and Internet protocol aspects

	Series Z
	Languages and general software aspects for telecommunication systems

	
	











_1063176175.unknown

_1063457181.unknown

_1063457521.unknown

_1063458327.unknown

_1064045751.unknown

_1065353858.unknown

_1063459190.unknown

_1063459293.unknown

_1063458239.unknown

_1063458260.unknown

_1063458266.unknown

_1063458245.unknown

_1063457747.unknown

_1063457492.unknown

_1063457504.unknown

_1063457471.unknown

_1063454384.unknown

_1063457153.unknown

_1063176372.unknown

_1063176853.unknown

_1063452963.unknown

_1063176703.unknown

_1063176294.unknown

_1063175605.unknown

_1063175942.unknown

_1063176092.unknown

_1063175690.unknown

_1063109887.unknown

_1063175503.unknown

_1063109401.unknown

